## ER1
ER1 FAILURE REPORTS
Tech: Viplav Dodeja
PN | SN | DATE | FAILURE REASON
CS-5146 | 1003253002 | 2026-09-16 | Pressure drops drastically during bench test. Averages 6 mmHg/min.
CS-5146 | 2204265018 | 2026-09-16 | Low QDM5 pressure during bench test (21.4 mmHg), and consistently failed console test. Tried different attempts at tuning, but failed each time. PID ranges from 25-15 and PID Integrals from 0.0150-0.0100.
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146
CS-5146

## ER2
ER2 FAILURE REPORTS
Tech: Viplav Dodeja
PN | SN | DATE | FAILURE REASON
CS-5046 | 2904263014 | 2026-09-06 | High vac at VDC. (1v=-46, 3v=-203, 5v=-359, 8v=-590). Multiple attempts at tuning yielded same results. Vac conistently higher than tolerance limit. At 0 VDC it is steady around -16.
CS-5046 | 1004265005 | 2026-09-16 | High pressure at VDC. (1v=-97, 3v=-248, 5v=-399, 8v=-622). Multiple attempts at tuning yielded same results. Vac conistently around 20 higher than tolerance limit. At 0 VDC it is steady around -46.
CS-5046 | "0611243006" | 2026-09-16 | Failed vac test. Pressure at 0 VDC holds steady at -16 mmHg. Low vac at VDC. (1v=-87, 3v=-228, 5v=-604, 8v=-687). Multiple attempts at tuning yielded same results. Vac conistently lower than tolerance limit.
CS-5046 | 1908243011 | 2026-09-16 | Failed vac test. Pressure at 0 VDC holds steady at -21 mmHg. Consistently low vac at VDC. (1v=-86, 3v=-238, 5v=-390, 8v=-638). Multiple attempts at tuning yielded same results. Vac conistently lower than tolerance limit.
CS-5046 | 2403253002 | 2026-09-16 | Failed vac test. Pressure at 0 VDC holds steady at -24 mmHg. Consistently low vac at VDC. (1v=-80, 3v=-234, 5v=-385, 8v=-606). Multiple attempts at tuning yielded same results. Vac conistently lower than tolerance limit.
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046
CS-5046